## Slide 1: Clinical BI Update
16-MAR-2022
Segun Jung, PhD
Associate Director of R&D and Clinical Bioinformatics

## Slide 2: Bioinformatics Team (AV)
Segun Jung, PhD
Associate Director
- Computational Biology
- NYU/UChicago
- Trail Running -> MBA
Christophe Magnan, PhD
Sr. Bioinformatics Scientist
- Computer Science
- UC Irvine
- Camping
Rudy Fabunan, MS
Bioinformatics Scientist
- Evolutionary Biology
- SDSU
- Running + dog
Hyunjun Nam, PhD
Bioinformatics Scientist
- Bioinformatics
- UCSD
- Tennis -> babysitting

## Slide 3: Introduction
Storage model
Clinical NGS pipelines (overview)
CNV
MSI
RNA expression
RNA fusion
Infrastructure
Challenges

## Slide 4: Storage Model
Isilon network
Total Size: 6.0P
Currently at 72% (03/15/22)
IT: NGS raw data archiving every day for older than 6 months or 1 year
R&D: purging for non-essential data (i.e., intermediate bam files) every day
Data transfer
Sequencing data
DNA/RNA Pipeline output folders

## Slide 5: NGS Pipeline - Demultiplexing
BCL (binary base call)
(36 hours)
Dumpltiplexing
Bcl2fastq v. 2.2
(~40min)
FASTQ
QIAseq323: 96
QIAseq297:192
IDT: 96
QIAseq with custom recipe (11 dark cycle)
Bclconvert v. 3.9
(~36min)
UMI processing and quality trimming
Fastp
(~15-20min)
UMI location
QIAseq: 12nt in R2
IDT: 3nt in both R1 and R2
Summary
QIAseq323 (DNA): ~100M reads
QIAseq297 (DNA/RNA): ~60M reads
IDT(RNA): ~30-60M reads
Total Time: ~ 1 hour
Email notification to the lab
Summary file
Batch PASS/FAIL: Q30 > 80%
Sample >5M
NTC <1M
Incorrect sample sheet
Detect incorrect sample sheet

## Slide 6: NGS Pipeline - Main
FASTQ
Alignment
(bwa mem)
UMI dedup
(fgbio)
Deduped Fastq
Ensemble Variant Calling
(vardict, mutect2, gatk4, platypus)
Post-processing
Total Read Depth (>100)
Alt Read Depth (>10)
AF: 3% for hotspot and 5% for non-hotspot
Artifacts
CPU assigned by 16 x number of FASTQ
~2hrs
~1hrs
BAM
VCF
DNA: hg19
Alignment
(bwa mem, Star)
UMI dedup
(umitools)
Deduped Fastq
Ensemble Variant Calling
(fusioncatcher, arriba, starfusion)
Post-processing
ML score filtering
Supporting reads
Artifacts
BAM
VCF
RNA: hg38
LIMS
(APvX)
VAaAST
QC
VCF (MSI)
CNV
QC
Parallelization whenever possible

## Slide 7: NGS Pipeline – Email notification

## Slide 8: Heme CNV - Overview
Dr. Christophe Magnan

## Slide 9: Heme CNV – Data Collection
Extracted Metadata Mapping, Merging,
And Filtering
Heme CNV Data
9,015 Concurrent Cases
364,210 Genes With CNV Information
From FISH
FISH Reports
387,938 Total
NGS Reports
113,456 Total
Isilon
257,951 BAM
176,547 VCF
FISH Results
Extraction
NGS Sample Metadata Extraction
NGS Data + Quality Controls

## Slide 10: Heme CNV – Features
QIAseq297
Pipeline
VCF
PureCN
CNVkit
Mosdepth
Segment-Level Log2(CN) x 4
Gene-Level Log2(CN) x 6
Normalized Coverage 10
…
BAM
Predictions From Existing Tools
PureCN and CNVkit selected
Baseline accuracy: 95% to 96%
Predicted Log2(CN) used as found
Gene and segment level results
Confidence intervals included
Expected > 0 for gain and < 0 for loss
Normalized Coverages
Classic and experimental features
11K samples to assess assay-specific coverage variations and biases
Comparison between gene coverage and other elements / expected normal
Expected > 1 for gain and < 1 for loss
Normalized Coverage 1
270 genes covered by the CNV methods
Dr. Christophe Magnan

## Slide 11: Heme CNV – Random Forest Classifier
252,858 genes with known CNV information
Random Forest
Training Algorithm
Random Forest CNV Model
Gene with unknown CNV information
Features +
Features
CNV Result
CNV Prediction
Loss: 71.23%
Gain: 03.56%
Normal: 25.21%
CNV Result

## Slide 12: Heme CNV – Pipeline Validation

## Slide 13: Heme CNV – Pipeline Validation Results
Model evaluated on each of the 62 genes
111,352 individual tests performed (1 gene x 1 sample)
Baseline accuracy (CNVkit, PureCN): 95% to 96%
Validation results summary:
Performance Characteristics | Percentage
Accuracy | 98.17%
Sensitivity | 95.98%
Specificity | 98.36%

## Slide 14: Heme CNV –Pipeline Validation Summary
Ensemble approach provides between 2% and 3% improvement
Limit of detection: 20% abnormal cells / tumor content
No limitation on the number of copies above the LOD
  Homozygous/heterozygous deletions both detected
  Gains detected with as low as 3 copies of the gene
No matched normal sample required
270 QIAseq297 genes covered by the methods
3 genes included in the recently launched ALL panel

## Slide 15: MSI for Solid Tumor – Training
MSI-H
(n=54)
MSS
(n=88)
Training dataset
Feature selection
(59 loci)
All samples in training/validation confirmed by Promega PCR assay or IHC
Baseline
(FFPE normal, MSS)
Panel | negative samples | positive samples
NeoTYPE™ Colorectal Tumor Profile | 30 | 17
NeoTYPE™ Lung Tumor Profile | 14 | 9
NeoTYPE™ Discovery Profile for Solid Tumors | 8 | 2
NeoTYPE™ Breast Tumor Profile | 5 | 1
NeoTYPE™ Pancreas Tumor Profile | 5 | 2
NeoTYPE™ Other Solid Tumor Profile | 4 | 1
NeoTYPE™ Endometrial Tumor Profile | 3 | 16
NeoTYPE™ Ovarian Tumor Profile | 3 | 1
NeoTYPE™ Precision Profile for Solid Tumors | 3 | 0
NeoTYPE™ Brain Tumor Profile | 2 | 1
NeoTYPE™ Esophageal Tumor Profile | 2 | 0
NeoTYPE™ Gastric Tumor Profile | 2 | 2
NeoTYPE™ GI Predictive Profile | 2 | 1
NeoTYPE™ Melanoma Profile | 2 | 0
NeoTYPE™ Cervical Tumor Profile | 1 | 1
NeoTYPE™ Head & Neck Tumor Profile | 1 | 0
NeoTYPE™ Liver/Biliary Tumor Profile | 1 | 0
Total | 88 | 54
Rudy Fabunan
CSV-SVPLR-0057-01

## Slide 16: MSI for Solid Tumor – 59 Markers
BAT40 | D3S1029 | D8S87 | KMT2D | GID4 | D20S100
AKT3 | ATR | D8S87 | POLE | D17S250 | EP300
MONO-27 | BAT25 | PRKDC | D13S175 | D17S787 | EP300
MSH2/BAT26 | D5S346 | TSC1 | HSP110-T17 | D18S35 | RBM10
MSH2 | CTNNA1 | D10S196 | D13S153 | TCF | AR
MSH6 | NPM1 | MRE11A | NR21 | D18S69 | AR
D2S123 | ARID1B | ATM | RAD51B | D18S64 | AR
XPO1 | PMS2 | ATM | RAD51B | D18S55 | STAG2
NR24 | D7S519 | NR22 | CREBBP | D18S61 | STAG2
TGFBR2 | D7S519 | KMT2D | BAT34C4 | TOP1 | 

## Slide 17: MSI for Solid Tumor –Val01
Prediction model
(59 loci)
All samples in training/validation confirmed by Promega PCR assay or IHC
TP | TN | FP | FN | Sen | Sep | Acc
47 | 339 | 5 | 2 | 96% | 99% | 98%
Val data
(393 samples)
True call
MSI-H
MSS
Indeterminate range: 0.45 <= x <= 0.5

## Slide 18: MSI for Solid Tumor –Val02
DNA input amount
Probability comparison for 100 ng vs 20 ng
@20ng – 4 sample QC fail due to average coverage < 500X

## Slide 19: MSI for Solid Tumor –Val03
Sample | True_call | MSI prediction | probability
A01-MSI01-20ng | MSI-H | MSS | 0.4
A02-MSI01-20ng | MSI-H | MSI-H | 0.73333333
A03-MSI01-20ng | MSI-H | MSI-H | 0.8
A04-MSI01-100ng | MSI-H | MSI-H | 0.9
A05-MSI01-100ng | MSI-H | MSI-H | 0.8
B01-MSI02-20ng | MSI-H | MSI-H | 0.73333333
B02-MSI02-20ng | MSI-H | MSI-H | 0.73333333
B03-MSI02-20ng | MSI-H | MSI-H | 0.6
B04-MSI02-100ng | MSI-H | MSI-H | 0.8
B05-MSI02-100ng | MSI-H | MSI-H | 0.53333333
C02-MSI03-20ng | MSI-H | MSI-H | 0.93333333
C03-MSI03-20ng | MSI-H | MSI-H | 0.73333333
C04-MSI03-100ng | MSI-H | MSI-H | 0.96666667
C05-MSI03-100ng | MSI-H | MSI-H | 0.86666667
@20ng & @20% - rare to receive in the lab
MSI04 & MSI05 omitted due to a sample issue

## Slide 20: RNA Solid Fusion
Strong fusion
25 supporting reads
Probability >0.5
Weak fusion
10 to 24 supporting reads
Probability >0.5
Require orthogonal confirmation
MET Exon 14 skipping & EGFRvIII mutation
Counting the chimeric junction sequence match
Positive: >50
Negative: < 10
Indeterminate: 10<x<50
CSV-SGFGS-0135_012621
Launched in January, 2021
Dr. Brad Thomas

## Slide 21: RNA Heme Fusion & Expression (ALL)
Fusion-positive
5 supporting reads
Probability >0.5
Expression
TPM values normalized by GUSB housekeeping gene
Cutoff established by Fei’s team
ALL launched
CSV-SGFGS-0179-01
Launched in February, 2022
Dr. Abhisek Ghosal

## Slide 22: Infrastructure
DNA/TNA
NovaSeq
(1) Support the clinical lab/operations in AV (and FM) to minimize TAT
(2) Support Laboratory Developed Test (LDT) development
Infrastructure
NGS Sequencing
BI QC
Data Analysis
Review & Reporting
Avg. 5 hrs
(Since Apr. 2021)
Avg. 10 hrs
(Oct. 2020 ─ Apr. 2021)
20 - 27 hrs
(Jul. 2019 ─ Oct. 2020)
?
128 cores
1144 cores
1808 cores
1808 cores
36 hrs
A few days?
Pipeline optimization
Avg. 5 hrs
Philip Torticill

## Slide 23
Head Node
R&D Cluster
2 compute nodes (184 CPUs)
Compute Node
node2
node1
Production Cluster in AV by the end of March
40 compute nodes (3728 CPUs)
CentOS 7/PCS cluster/Torque 6.0.1/Maui
List of pipelines
Heme/Solid DNA pipelines
Heme/Solid RNA pipelines
Heme/Solid RNA QC pipelines
Additional offering
NGS only Solid Lung 2.0 launch
Comprehensive/Discovery launch
Solid RNA Volume increases
Heme ALL
Issues
In general, 3-4 batches of Heme/Solid in the cluster (96 samples per batch) instead of 2 batches
Only two-three batches, but Solid RNA batches
Deploy the pipelines to other locations
Improvement
Increase # of CPUs per sample (i.e. 16 -> 24)
Improve in alignment with bwa-mem2
GPU
Head node
Compute node
node2
node1
node3
node39
node40
….
Infrastructure

## Slide 24: Challenges
Staffing – current opening
Sr. Bioinformatics Scientist (Christophe’s replacement)
Bioinformatics Scientist (Hyunjun’s replacement)
Cloud Bioinformatics Engineer (on hold until Q3 per HR)
Hard to get good bioinformatician (quality vs. compensation)
Scalability
Need to implement a cloud option in case the volume in the lab increases

## Slide 25: NGS Pipeline - Demultiplexing
Email notification examples

## Slide 26: Pipeline Validation Results (2)
Highest concordance when comparing NGS and FISH at the exact same location
Gene | Direction | Positive Cases |  |  | Negative Cases |  |  | All Cases |  | 
 |  | Total | Conc. | Sensitivity | Total | Conc. | Specificity | Total | Conc. | Accuracy
ATM | loss | 72 | 72 | 100.00% | 641 | 629 | 98.13% | 713 | 701 | 98.32%
CBFB | loss | 23 | 22 | 95.65% | 517 | 506 | 97.87% | 540 | 528 | 97.78%
EGR1 | loss | 171 | 169 | 98.83% | 1,541 | 1,538 | 99.81% | 1,712 | 1,707 | 99.71%
KMT2A | gain | 27 | 25 | 92.59% | 474 | 472 | 99.58% | 501 | 497 | 99.20%
MET | loss | 113 | 113 | 100.00% | 1,593 | 1,582 | 99.31% | 1,706 | 1,695 | 99.36%
NF1 | loss | 15 | 15 | 100.00% | 908 | 908 | 100.00% | 923 | 923 | 100.00%
TERT | loss | 10 | 10 | 100.00% | 1,723 | 1,709 | 99.19% | 1,733 | 1,719 | 99.19%
TP53 | loss | 76 | 73 | 96.05% | 1,567 | 1,556 | 99.30% | 1,643 | 1,629 | 99.15%

## Slide 27: MSI Prediction Scheme
MSI-H
(n=93)
MSS
(n=157)
Training dataset
MSS
MSI-H
Logistic regression with LBFGS solver
Prediction model
Clinical Samples
Blind dataset
?
95% probability
95% probability
95% probability
MSI-H
MSS
REFLEX
msisensor
msisensor

## Slide 28: Features -- Peaks

## Slide 29: Solid CNV – Summary
Normal dataset established aligned with the QC in current validation report
  QC criteria
    Minimum 5M reads and 500X avg. coverage
    Uniformity: 95%
    100X: 98%
  Pool of Normas: 15 TNA
Retrospective cases collected and data processed
  A single cutoff
    0.4 for amplification (except for ERBB2 at 0.2) and -0.35 deletion
gene | Log2 cutoff | Dataset |  |  |  | Performance |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  | training (60%) |  | validation (40%) |  | training (60%) |  |  |  |  |  |  | validation (40%) |  |  |  |  |  |  | 
 |  | positive | negative | positive | negative | TP | TN | FP | FN | Sen | Spe | Acc | TP | TN | FP | FN | Sen | Spe | Acc | Acc
MET | 0.4 | 115 | 4722 | 41 | 2033 | 100 | 4321 | 431 | 15 | 0.87 | 0.91 | 0.91 | 35 | 1866 | 177 | 6 | 0.85 | 0.91 | 0.91 | 0.91
ERBB2 | 0.2 | 126 | 1687 | 82 | 1128 | 87 | 1593 | 106 | 39 | 0.69 | 0.94 | 0.92 | 64 | 1054 | 80 | 20 | 0.76 | 0.93 | 0.92 | 0.92
EGFR | 0.4 | 14 | 138 | 3 | 99 | 13 | 118 | 20 | 1 | 0.93 | 0.86 | 0.86 | 2 | 86 | 13 | 0 | 1 | 0.87 | 0.87 | 0.87
PTEN | -0.35 | 70 | 1639 | 47 | 1092 | 42 | 1387 | 252 | 28 | 0.6 | 0.85 | 0.84 | 27 | 913 | 179 | 20 | 0.57 | 0.84 | 0.83 | 0.83
panel of normals (TNA: N =15)