## Slide 1: Heme CNV - Overview
Dr. Christophe Magnan

## Slide 2: Heme CNV – Data Collection
Extracted Metadata Mapping, Merging,
And Filtering
Heme CNV Data
9,015 Concurrent Cases
364,210 Genes With CNV Information
From FISH
FISH Reports
387,938 Total
NGS Reports
113,456 Total
Isilon
257,951 BAM
176,547 VCF
FISH Results
Extraction
NGS Sample Metadata Extraction
NGS Data + Quality Controls

## Slide 3: Heme CNV – Features
QIAseq297
Pipeline
VCF
PureCN
CNVkit
Mosdepth
Segment-Level Log2(CN) x 4
Gene-Level Log2(CN) x 6
Normalized Coverage 10
…
BAM
Predictions From Existing Tools
PureCN and CNVkit selected
Baseline accuracy: 95% to 96%
Predicted Log2(CN) used as found
Gene and segment level results
Confidence intervals included
Expected > 0 for amp and < 0 for del
Normalized Coverages
Classic and experimental features
11K samples to assess assay-specific coverage variations and biases
Comparison between gene coverage and other elements / expected normal
Expected > 1 for amp and < 1 for del
Normalized Coverage 1
270 genes covered by the CNV methods
Dr. Christophe Magnan

## Slide 4: Heme CNV – Random Forest Classifier
252,858 genes with known CNV information
Random Forest
Training Algorithm
Random Forest CNV Model
Gene with unknown CNV information
Features +
Features
CNV Result
CNV Prediction
Loss: 71.23%
Gain: 03.56%
Normal: 25.21%
CNV Result

## Slide 5: Heme CNV – Pipeline Validation

## Slide 6: Heme CNV – Pipeline Validation Results
Model evaluated on each of the 62 genes
111,352 individual tests performed (1 gene x 1 sample)
Baseline accuracy (CNVkit, PureCN): 95% to 96%
Validation results summary:
Performance Characteristics | Percentage
Accuracy | 98.17%
Sensitivity | 95.98%
Specificity | 98.36%

## Slide 7: Heme CNV –Pipeline Validation Summary
Ensemble approach provides between 2% and 3% improvement
Limit of detection: 20% abnormal cells / tumor content
No limitation on the number of copies above the LOD
  Homozygous/heterozygous deletions both detected
  Gains detected with as low as 3 copies of the gene
No matched normal sample required
270 QIAseq297 genes covered by the methods
3 genes included in the recently launched ALL panel